## Slide 1: Vectorized Coordinate Descent for CT Reconstruction
Greg Buzzard (joint with Charlie Bouman)
Department of Mathematics
Purdue University

## Slide 2: Outline
Overview of CT imaging and MBIR reconstruction
  30 years ago: research in CT is dead
  Applications: Medical, industrial, 3d printing, neutron, security
Challenges with CT:
  Long-range interactions – a large, dense matrix with a long tail of singular values – difficult to invert or do gradient descent
  Gradient descent gets low spatial frequencies quickly but stalls on high frequencies
  ICD gets high spatial frequencies but is difficult to parallelize
Key ideas of VCD:
  Unify gradient descent and ICD as two ends of a spectrum of updates on subsets
  Start with gradient descent to get low frequencies. Then update on smaller subsets to get higher frequencies.
  Subsampling improves conditioning and simplifies the gradient structure
  Diagonally preconditioned gradient steps for small subsets are nearly optimal solves.
Implementation in software

## Slide 3: CT Overview and Model Based Iterative Reconstruction
30 years ago: research in CT is dead

## Slide 4: Computed Tomographic (CT) Imaging
Parallel Beam CT: synchrotrons, electron microscopy, nano-X-ray sources
Fan Beam CT: Industrial CT
Cone Beam CT: Industrial CT, C-arm Scanners
Multi-Slice Helical CT: Medical, transportation security

## Slide 5: CT Forward Model
Measurements
System Matrix
Volume to be Reconstructed
Noise
Ongoing Problems:
    Not enough measurements: sparse or missing views, etc.
    Low quality data: high noise, low dosage, short exposure, etc.
    Model mismatch: metal, beam-hardening, scatter, poly-energetic, etc.
    Resolution loss: detector blur, motion blur, X-ray spot size, etc.
    Direct reconstruction (FBP, FDK) gives poor results in these cases.
Applications:
    Medical, scientific, industrial, and security

## Slide 6: Model-Based Iterative Reconstruction (MBIR) for CT
Physical system
Image Model: p(x)
Forward model
Difference
Error Sinogram ()
Fwd Model
Adjust to reduce the loss
Measured data
Target object
Estimated recon
data fitting + prior

## Slide 7: MBIR for 64 slice GE VCT Scanner circa 2011
Previous state-of-the-art 3D Recon
GE MBIR
Purdue/Notre Dame/GE algorithm
MBIR-CT Advantages:
    Can model the physical system and measurement noise properties
    Can use prior information (now including NN priors via PnP)
    Can modify to make corrections for beam hardening and other artifacts

## Slide 8: MBIR-CT Challenges
Increasing scan sizes and volumes:
  1K x 1K sensor 1K x 1K x 1K volume: 4GB
  2K x 2K sensor 2K x 2K x 2K volume: 32GB
Computation time:
  Each iteration makes an update to each voxel – would like to parallelize by voxel.
  Each voxel depends on all the views – nonlocal memory access.
  Need to use very few iterations to limit reconstruction times.
View
Detector

## Slide 9: MBIR-CT Challenges
Increasing scan sizes and volumes:
  1K x 1K sensor 1K x 1K x 1K volume: 4GB
  2K x 2K sensor 2K x 2K x 2K volume: 32GB
Computation time:
  Each iteration makes an update to each voxel – would like to parallelize by voxel.
  Each voxel depends on all the views – nonlocal memory access.
  Need to use very few iterations to limit reconstruction times.
Approaches
  Gradient descent: Slow to converge at high spatial frequencies – stalls out after a few iterations.
  Iterative coordinate descent (optimal update of one voxel at a time): Difficult to parallelize.
View
Detector
 | Easy to parallelize | Fast convergence
Gradient descent | ✅ | ❌
ICD | ❌ | ✅

## Slide 10: Why don’t people use MBIR?
It’s too slow
  20 to 200x slower than FBP
It’s too hard to use
  Difficult to pick parameters
It’s too difficult to implement
  Need to “roll your own” for every new geometry
This is not Frictionless Reproducibility!![1]
  The combination of Data Sharing, Code Sharing, and Challenges has dramatically accelerated research in the AI field.
[1] David Donoho, “Data Science at the Singularity”, Harvard Data Science Review, 6(1), Jan. 2024.

## Slide 11: MBIRJAX: A New Approach
VCD (Vectorized Coordinate Descent):
    A hybrid of Gradient Descent (GD) and Iterative Coordinate Descent (ICD)
    Fast convergence and Fast implementation
Hardware Structure
Fast and large
Scalable
Software Structure
Expressive and parallel
Portable and efficient
Algorithmic Structure
Fast convergence
Implementable
GPU, CPU, TPU
JAX, PyTorch, Julia
VCD, GD, ICD
our focus today

## Slide 12: MBIRJAX Software Package
General purpose reconstruction software
    Open-source BSD 3-clause
    Easy to use and install with pip
    Easy to build new geometries:
      Parallel beam
      Cone beam
      More to come
URL:
    https://github.com/cabouman/mbirjax
    https://mbirjax.readthedocs.io

## Slide 13: Important Points to Remember
Simplified problem:
Gradient Descent (GD):
  ✔ Fast updates
  ❌ Slow convergence
Coordinate Descent (CD):
  ✔ Fast convergence
  ❌ Slow updates
Is there a better solution?

## Slide 14: Coordinate Descent
    Update each pixel in sequence
    Keep a state vector for efficiency
loss
slope
curvature
starting point
Newton Rapson Update
def:
# Initialize state
# Precompute
for each iteration:
for each pixel :
# Compute pixel update
#Update state
Return
Not practical on modern GPU hardware!

## Slide 15: Vectorized Coordinate Descent
Partition pixels into subsets (colors)
For each color:
    Update pixels in parallel
    Use coordinate descent optimization
Cycle through subsets in sequence
def:
# Initialize
for each iteration:
for each subset :
# Compute pixel updates
#Update state
Return
Bugzilla!
-Algorithmic instability that occurs when too many pixels are updated simultaneously.
-Conventional solution is to damp updates by a factor of [1,2].
[1] J. A. Fessler, E. P. Ficaro, N. H. Clinthorne, and K. Lange, “Grouped-co-ordinate ascent algorithms for penalized-likelihood transmission image reconstruction,” IEEE Trans. Med. Imag., vol. 16, pp. 166–175, Apr. 1997.
[2] K. D. Sauer, S. Borman, and C. A. Bouman, “Parallel Computation of Sequential Pixel Updates in Statistical Tomographic Reconstruction,” vol. 2, pp. 93-96, IEEE Int'l Conf. on Image Proc., Washington, DC, Oct. 22-25, 1995.

## Slide 16: Bugzilla Swatter
Always Converges 
def:
# Initialize
for each iteration:
for each subset :
# Compute pixel updates
#Update state
Return
def:
# Initialize
for each iteration:
for each subset :
# Compute image change
    # Compute optimal step
    # Update state
Return
Unstable 
Bugzilla Swatter

## Slide 17: Bugzilla Swatter
Always Converges 
Bugzilla Swatter
Unstable 

## Slide 18: The "Aha!" moment!
Conversation with self:
    Oh, so you can do coordinate descent (CD) at every pixel simultaneously??
    But if you update every pixel, then that’s not coordinate descent, right??
    Oh, if you update every pixel, I guess it must be gradient descent…
Conclusion: VCD is both GD and CD!

## Slide 19: Multi-Granular VCD
Change the number of subsets during optimization
Vary between GD and CD and everything in between
A little like SGD, but the batches are in the domain
courser granularity
Gradient Descent
Better low frequency convergence
finer granularity
Coordinate Descent
Better high frequency convergence

## Slide 20: Unifying Method: Vectorized Coordinate Descent
 | Easy to parallelize | Fast convergence
Gradient descent | ✅ | ❌
ICD | ❌ | ✅
VCD | ✅ | ✅
VCD outline:
  Each iteration (one partition):
    Partition all voxels into multiple subsets, each of the same size.
    Update each subset using diagonally preconditioned gradient descent and an optimal step size.
  Loop over partitions:
    Start with all voxels in a single subset (gradient descent).
    Reduce subset size with each iteration, but not all the way to single voxels.
This gives
  Broad outlines of recon in first few iterations – like gradient descent.
  Sharp edges after just a few iterations – like ICD.
  Easy to parallelize.

## Slide 21: VCD Performance
Gradient descent
ICD
VCD gets the smooth interior and sharp edges
10 iterations total
VCD starts like gradient descent but ends like ICD

## Slide 22: Convergence: Recons
Gradient descent
3 iterations
5 iterations
10 iterations
ICD
VCD
Phantom

## Slide 23: Convergence: |error|
Gradient descent
ICD
VCD
3 iterations
5 iterations
10 iterations
Phantom

## Slide 24: Convergence: |FFT(error)| in dB
Gradient descent
ICD
VCD
3 iterations
5 iterations
10 iterations
|FFT(Phantom)|

## Slide 25: Why does VCD work? Initial Observations
Goal: solve
Assume is quadratic with nearest neighbor coupling.
Gradient:

## Slide 26: Why does VCD work? Initial Observations
Observation 1: is long-range blur
Approximate with a convolution:
where is a 2D point spread function
Leads to long range interactions and a dense matrix.
Goal: solve
Assume is quadratic with nearest neighbor coupling.
Gradient:
Observation 2: amplifies low spatial frequencies
In the Fourier domain:
where is the 2D DTFT of :
Very large for near the origin (low frequencies).
PSF : heat map
PSF : restricted to 1D

## Slide 27: Analysis of gradient descent
Observation 1: is long-range blur
Goal: solve
Observation 2: amplifies low spatial frequencies
Observation 1 means is dense and not close to diagonal
  Hard to invert or even get a good approximate inverse.
Observations 2 and 3 mean gradient descent must take small steps, even if we choose an optimal step size.
Gradient descent:
In Fourier space:
For low frequencies:
Observation 3: for low spatial freqs, in Fourier domain
|FFT(phantom)| in dB
for (fast low freq convergence)
means
for large (slow high freq convergence)
Updates are concentrated on low frequencies.
|FFT|

## Slide 28: Analysis of coordinate descent
Goal: solve
|FFT(phantom)| in dB
Observation 4: If we fix all but one element of , we can solve for that element exactly.
Iterating through the coordinates gives ICD.
Changing a single pixel affects all the Fourier modes, high frequencies the most.
Cycling through the pixels gives ICD.
VCD uses a partition of sparse subsets of pixels for parallelism.
|FFT|
A few iterations of ICD
Reduces the low frequencies
Introduces new high frequencies, which decay slowly

## Slide 29: Analysis of VCD: Changing subset size
Goal: solve
|FFT(phantom)| in dB
|FFT|
Observation 5: Gradient descent is a good start
Most images have more energy in low frequencies than high frequencies. One or two steps of gradient descent capture this.
Observation 6: Subsampling in space makes each subproblem easier
1. becomes closer to diagonal: subsampling by yields a psf
2. becomes completely diagonal when there are no nearest neighbors.
So, for large subsampling factor we can solve each subproblem almost exactly using diagonal preconditioning and optimal step size.
Observation 7: Subsampling in space spreads each update more evenly across multiple frequencies in Fourier space
The diagonal matrix in Observation 6 corresponds to multiplication by a constant in Fourier space:
But that is in Fourier space of the subsampled image!
Even a constant update to a subsampled image updates many frequencies in the Fourier space of the full image.
Updates to both low and high frequencies.

## Slide 30: Analysis of VCD: Changing subset size
Goal: solve
|FFT(phantom)| in dB
|FFT|
Observation 5: Gradient descent is a good start
Most images have more energy in low frequencies than high frequencies. One or two steps of gradient descent capture this.
Observation 6: Subsampling in space makes each subproblem easier
1. becomes closer to diagonal: subsampling by yields a psf
2. becomes completely diagonal when there are no nearest neighbors.
So, for large subsampling factor we can solve each subproblem almost exactly using diagonal preconditioning and optimal step size.
Observation 7: Subsampling in space spreads each update more evenly across multiple frequencies in Fourier space
The diagonal matrix in Observation 6 corresponds to multiplication by a constant in Fourier space:
But that is in Fourier space of the subsampled image!
Even a constant update to a subsampled image updates many frequencies in the Fourier space of the full image.
Updates to both low and high frequencies.

## Slide 31: Analysis of VCD: Changing subset size
Goal: solve
|FFT(phantom)| in dB
|FFT|
Observation 5: Gradient descent is a good start
Most images have more energy in low frequencies than high frequencies. One or two steps of gradient descent capture this.
Observation 6: Subsampling in space makes each subproblem easier
1. becomes closer to diagonal: subsampling by yields a psf
2. becomes completely diagonal when there are no nearest neighbors.
So, for large subsampling factor we can solve each subproblem almost exactly using diagonal preconditioning and optimal step size.
Observation 7: Subsampling in space spreads each update more evenly across multiple frequencies in Fourier space
The diagonal matrix in Observation 6 corresponds to multiplication by a constant in Fourier space:
But that is in Fourier space of the subsampled image!
Even a constant update to a subsampled image updates many frequencies in the Fourier space of the full image.
Updates to both low and high frequencies.

## Slide 32: Convergence (FFT error in dB)
Gradient descent
ICD
VCD
3 iterations
5 iterations
10 iterations
|FFT(Phantom)|

## Slide 33: VCD Key Ideas
Smaller subsets make the system for each subproblem closer to diagonal.
  Full problem is stiff and dense – difficult to solve exactly or even take a large step.
  Can solve each smaller subproblem almost exactly.
Smaller subsets blend the high and low frequency Fourier modes
  These low frequency Fourier modes are associated with the largest singular values in the full problem, so this reduces the stiffness.
Start with gradient descent (all pixels) and move towards ICD (few pixels)
  It’s helpful to get the low frequency Fourier modes close to accurate before going to subsets.
Point spread function falls off like 1/distance.
Leads to long range interactions and a dense matrix.
Use subsets to make the problem easier.

## Slide 34: MBIRJAX Software Package
General purpose reconstruction software
    Open-source BSD 3-clause
    Easy to use and install with pip
    Easy to build new geometries:
      Parallel beam
      Cone beam
      More to come
URL:
    https://github.com/cabouman/mbirjax
    https://mbirjax.readthedocs.io

## Slide 35: Initialization with FBP
Gradient descent
ICD
10 iterations total

## Slide 36: Recon parameters
VCD parameters
param_dict = { # Set parameters 'num_views' : 128, 'num_det_rows' : 40, 'num_det_channels' : 256, 'start_angle' : 0, 'end_angle' : np.pi, 'granularity': [1, 8, 48, 128, 128,], 'partition_sequence' : [0, 1, 2, 3, 3, 3, 3, 3, 3, 3, 3, 3,], 'num_iterations' : 10 }
Gradient Descent parameters: same but using 1 subset for each iteration
ICD parameters: same but using 128 subsets for each iteration